## Summary
Scanner/Coil | Philips 3.0T Achieva / 16 ch breast coil
Software Version | 5.1.7.1 (plus custom software patch; '+' indicates non-standard parameters included in the patch)
Sequence | Name | Scan duration (mm:ss)
DWI | DWI_EPI_b0200800 | 0:04
B1 | Bloch-Siegert_FA300_FH_wR3Param | 0:02
B0 | B0 3D InPhase | 0:00
qMT | qMT 3D Pulsed Trig | 0:02
VFA | multi-flip_T1-map | 0:01
DCE | dynamic_smartTX_6deg | 0:06
T1w Anatomical | THRIVE SENSE | 0:02

## DWI
DWI_EPI_b0200800 03:57.0
INFO PAGE |  | GEOMETRY |  | CONTRAST
Total scan duration | 03:57.0 | Nucleus | H1 | Scan type | Imaging
Rel. SNR | 1 | Uniformity | Body-tuned | Scan mode | MS
Act. TR (ms) | 3000 | FOV FH (mm) | 192 | technique | SE
Act. TE (ms) | 43 | AP (mm) | 192 | Modified SE | no
ACQ matrix M x P | 96 x 96 | RL (mm) | 50 | +REF/Crush Control | no
ACQ voxel MPS (mm) | 2.00 / 2.00 / 5.00 | Voxel size FH (mm) | 2 | Acquisition mode | cartesian
REC voxel MPS (mm) | 1.33 / 1.33 / 5.00 | AP (mm) | 2 | Fast Imaging mode | EPI
Scan percentage (%) | 100 | Slice thickness (mm) | 5 | shot mode | single-shot
Packages | 1 | Recon voxel size (mm) | 1.333333 | Echoes | 1
Min. slice gap (mm) | 0 | Small FOV imaging | no | partial echo | no
Diff sep/dur/spe/str/max | 20.3/10.2/0.9/79.8/93.0 | Fold-over suppression | rest | TE | user defined
EPI factor | 39 | slab thickness | half FOV | (ms) | 43
WFS (pix) / BW (Hz) | 12.186 / 35.7 | Reconstruction matrix | 144 | Flip angle (deg) | 90
BW in EPI freq. dir. (Hz) | 2963.6 | SENSE | yes | TR | user defined
+BW EX (Hz) | 565.0 | P reduction (AP) | 3 | (ms) | 3000
+BW REF (Hz) | 735.0 | k-t BLAST | no | Halfscan | yes
+GR EX (mT/m) | 2.45 | Stacks | 1 | factor | 0.6285715
+GR REF (mT/m) | -2.94 | type | parallel | Water-fat shift | minimum
+Null-Freq (Hz) | 321.1 | slices | 10 | RF Shims | smart
Min. TR (ms) | 1804 | slice gap | user defined | Shim | IB-smart
SPAIR offset act./default (Hz) | 220 [220] | gap (mm) | 0 | mDIXON | no
SAR / local torso | < 43 % | slice orientation | sagittal | Fat suppression | SPAIR
Whole body / level | < 0.7 W/kg / normal | fold-over direction | AP | power | 2
SED | < 0.2 kJ/kg | fat shift direction | P | inversion delay | auto
B1+rms / Coil Power | 1.50 uT / 41 % | Minimum number of packages | 1 | suppr. level | strong
Max B1+rms | 1.50 uT | Slice scan order | default | frequency offset | user defined
PNS / level | 70 % / normal | PlanAlign | no | offset (Hz) | 220
dB/dt | 54.3 T/s | REST slabs | 0 | Grad Rev Fat suppression | yes
Sound Pressure Level (dB) | 13.53854 | Catheter tracking | no | +Mode | default
max d¦B¦/dt (epi) [Tesla/second] | 54.3 (29.8) | Interactive positioning | no | Water suppression | no
+Ref Pulse Shape |  | External control | no | BB pulse | no
+Ref Pulse Dur [msec] |  | Allow table movement | no | MTC | no
+Ref MAX Depahse |  | OFFC/ANG |  | Research prepulse | no
+Ref MIN Dephase |  | Stacks | 1 | Diffusion mode | DWI
+Crusher b value |  | Stack Offc. AP (P=+mm) | -25 | sequence | SE
MOTION |  | RL (L=+mm) | -100 | gradient duration | maximum
Cardiac synchronization | no | FH (H=+mm) | 25 | gradient overplus | yes
Heart rate > 250 bpm | no | Ang. AP (deg) | 0 | +Separation | no
Respiratory compensation | no | RL (deg) | 0 | +OGSE | no
Navigator respiratory comp | no | FH (deg) | 0 | nr of b-factors | 3
Flow compensation | no | Free rotatable | no | b-factor order | user defined
Temporal slice spacing | default |  |  | b-factors | 0, 200, 800, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0
NSA | 6 |  |  | average high b | user defined
SMART | no |  |  | b-factor averages | (0) 1, (200) 1, (800) 3, , , , , , , , , , , , , , , , , , , , , , , , , , , , ,
DYN/ANG |  |  |  | Elastography mode | no
Manual start | no |  |  | Multi-transmit | yes
+Abuse dynamic loop | no |  |  | Transmit channels | both
Dynamic study | no |  |  | SAR mode | high
Arterial Spin labeling | no |  |  | B1 mode | default
POST/PROC |  |  |  | SAR Patient data | auto
Preparation phases | full |  |  | PNS mode | moderate
Interactive F0 | select |  |  | Gradient mode | enhanced
SmartPlan survey | no |  |  | SofTone mode | no
+Whole Body Fat-Water | no
B0 field map | no
B1 field map | no
MIP/MPR | no
Images | M, no, no, no
Autoview image | M
Calculated images | no, no, no, no
Reference tissue | Breast
EPI 2D phase correction | no
Recon compression | No
Preset window contrast | soft
Reconstruction mode | immediate
Save raw data | no
Hardcopy protocol | no
Image filter | system default
Geometry correction | default

## B1
Bloch-Siegert_FA300_FH_wR3Param 01:44.1
INFO PAGE |  | GEOMETRY |  | CONTRAST
Total scan duration | 01:44.1 | Nucleus | H1 | Scan type | Imaging
Rel. SNR | 1 | Uniformity | Body-tuned | Scan mode | MS
Act. TR/TE (ms) | 491 / 5.4 | FOV AP (mm) | 256 | technique | FFE
ACQ matrix M x P | 104 x 102 | FH (mm) | 256 | Contrast enhancement | no
ACQ voxel MPS (mm) | 2.46 / 2.51 / 5.00 | RL (mm) | 50 | Acquisition mode | cartesian
REC voxel MPS (mm) | 1.33 / 1.33 / 5.00 | Voxel size AP (mm) | 2.5 | Fast Imaging mode | none
Scan percentage (%) | 98.07692 | FH (mm) | 2.509804 | Echoes | 1
Packages | 1 | Slice thickness (mm) | 5 | partial echo | no
Min. slice gap (mm) | 0 | Recon voxel size (mm) | 1.333333 | shifted echo | no
Act. WFS (pix) / BW (Hz) | 0.151 / 2876.0 | Fold-over suppression | no | TE | shortest
Min. WFS (pix) / Max. BW (Hz) | 0.151 / 2876.0 | Reconstruction matrix | 192 | Flip angle (deg) | 90
SAR / local torso | < 100 % | SENSE | no | TR | shortest
Whole body / level | < 1.7 W/kg / normal | k-t BLAST | no | Halfscan | no
SED | < 0.1 kJ/kg | Stacks | 1 | Water-fat shift | minimum
B1+rms / Coil Power | 2.29 uT / 96 % | type | parallel | RF Shims | smart
Max B1+rms | 2.29 uT | slices | 10 | Shim | IB-smart
+Max B1+ of BS pulse | 9.771 uT | slice gap | user defined | mDIXON | no
PNS / level | 80 % / normal | gap (mm) | 0 | Fat suppression | no
dB/dt | 96.0 T/s | slice orientation | sagittal | Water suppression | no
Sound Pressure Level (dB) | 16.81488 | fold-over direction | FH | MTC | no
max d¦B¦/dt (epi) [Tesla/second] | 96.0 (0.0) | fat shift direction | P | Research prepulse | no
MOTION |  | Minimum number of packages | 1 | Diffusion mode | no
Cardiac synchronization | no | Slice scan order | default | Elastography mode | no
Heart rate > 250 bpm | no | PlanAlign | no | Multi-transmit | yes
Respiratory compensation | no | REST slabs | 0 | Transmit channels | both
Navigator respiratory comp | no | Catheter tracking | no | SAR mode | high
Flow compensation | no | Interactive positioning | no | B1 mode | default
Temporal slice spacing | default | External control | no | SAR Patient data | auto
fMRI echo stabilisation | no | Allow table movement | no | PNS mode | moderate
NSA | 1 | OFFC/ANG |  | Gradient mode | maximum
DYN/ANG |  | Stacks | 1 | SofTone mode | no
Angio / Contrast enh. | no | Stack Offc. AP (P=+mm) | -25
Quantitative flow | no | RL (L=+mm) | -100
Manual start | no | FH (H=+mm) | 25
+Abuse dynamic loop | no | Ang. AP (deg) | 0
Dynamic study | no | RL (deg) | 0
Arterial Spin labeling | no | FH (deg) | 0
POST/PROC |  | Free rotatable | no
Preparation phases | full
Interactive F0 | no
SmartPlan survey | no
+Whole Body Fat-Water | no
B0 field map | no
B1 field map | yes
field map technique | +phase shift
+pulse choice | optimized
+flip angle | 300
+spoiler dur (ms) | 0
MIP/MPR | no
SWIp | no
Images | M, R, I, no
Autoview image | M
Calculated images | B1, no, no, no
Reference tissue | Breast
Recon compression | No
Preset window contrast | soft
Reconstruction mode | real time
Save raw data | no
Hardcopy protocol | no
Image filter | system default
Geometry correction | default

## B0
B0 3D InPhase 00:09.0
INFO PAGE |  | GEOMETRY |  | CONTRAST
Total scan duration | 00:09.0 | Nucleus | H1 | Scan type | Imaging
Rel. SNR | 1 | Uniformity | Classic | Scan mode | 3D
Act. TR/TE (ms) | 12 / 4.6 | FOV AP (mm) | 256 | technique | FFE
ACQ matrix M x P | 84 x 85 | FH (mm) | 256 | loop order | zy_order
ACQ voxel MPS (mm) | 3.05 / 3.01 / 10.0 | RL (mm) | 50 | Contrast enhancement | T1
REC voxel MPS (mm) | 1.33 / 1.33 / 5.00 | Voxel size AP (mm) | 3 | Acquisition mode | cartesian
Scan percentage (%) | 101.1905 | FH (mm) | 3 | Fast Imaging mode | none
Act. WFS (pix) / BW (Hz) | 0.300 / 1446.5 | RL (mm) | 5 | 3D non-selective | no
Min. WFS (pix) / Max. BW (Hz) | 0.279 / 1554.7 | Recon voxel size (mm) | 1.333333 | Echoes | 1
Min. TR/TE (ms) | 9.0 / 4.4 | Fold-over suppression | no | partial echo | no
SAR / local torso | < 5 % | Slice oversampling | default | shifted echo | no
Whole body / level | < 0.1 W/kg / normal | RF select. FOS | no | TE | in-phase
SED | 0.0 kJ/kg | Reconstruction matrix | 192 | (ms) | 4.603275
B1+rms / Coil Power | 0.49 uT / 4 % | SENSE | no | Flip angle (deg) | 8
Max B1+rms | 0.49 uT | k-t BLAST | no | TR | user defined
PNS / level | 24 % / normal | Overcontiguous slices | yes | (ms) | 12
dB/dt | 12.8 T/s | Stacks | 1 | Halfscan | no
Sound Pressure Level (dB) | 13.10469 | slices | 10 | Water-fat shift | minimum
max d¦B¦/dt (epi) [Tesla/second] | 12.8 (0.0) | slice orientation | sagittal | RF Shims | smart
MOTION |  | fold-over direction | FH | Shim | IB-smart
Cardiac synchronization | no | fat shift direction | P | mDIXON | no
Heart rate > 250 bpm | no | Multi-chunk | no | Fat suppression | no
Respiratory compensation | no | PlanAlign | no | Water suppression | no
Navigator respiratory comp | no | REST slabs | 0 | MTC | no
Flow compensation | yes | Catheter tracking | no | Research prepulse | no
fMRI echo stabilisation | no | Interactive positioning | no | Diffusion mode | no
NSA | 1 | External control | no | Elastography mode | no
DYN/ANG |  | Allow table movement | no | Multi-transmit | yes
Angio / Contrast enh. | no | OFFC/ANG |  | Transmit channels | both
Quantitative flow | no | Stacks | 1 | SAR mode | high
Manual start | no | Stack Offc. AP (P=+mm) | -25 | B1 mode | default
+Abuse dynamic loop | no | RL (L=+mm) | -100 | SAR Patient data | auto
Dynamic study | no | FH (H=+mm) | 25 | PNS mode | high
Arterial Spin labeling | no | Ang. AP (deg) | 0 | Gradient mode | default
POST/PROC |  | RL (deg) | 0 | SofTone mode | no
Preparation phases | +sameprep | FH (deg) | 0
Interactive F0 | no | Free rotatable | no
SmartPlan survey | no
+Whole Body Fat-Water | no
B0 field map | yes
delta TE (ms) | 2.3
B1 field map | no
MIP/MPR | no
SWIp | no
Images | M, R, I, P
Autoview image | M
Calculated images | B0, no, no, no
Reference tissue | Skeletal muscle
Preset window contrast | soft
Reconstruction mode | real time
Save raw data | no
Hardcopy protocol | no
Image filter | system default
Geometry correction | default
Elliptical k-space shutter | default

## qMT
qMT 3D Pulsed Trig 01:38.4
INFO PAGE |  | GEOMETRY |  | CONTRAST
Research Options used | 22 | Nucleus | H1 | Scan type | Imaging
Total scan duration | 01:38.4 | Uniformity | Body-tuned | Scan mode | 3D
Rel. SNR | 1 | FOV AP (mm) | 256 | technique | FFE
Act. TR/TE (ms) | 48 / 6.4 | FH (mm) | 256 | Contrast enhancement | T1
Dyn. scan time | 00:10.9 | RL (mm) | 50 | Acquisition mode | cartesian
Time to k0 | 00:09.0 | Voxel size AP (mm) | 2 | Fast Imaging mode | EPI
ACQ matrix M x P | 128 x 126 | FH (mm) | 2 | 3D non-selective | no
ACQ voxel MPS (mm) | 2.00 / 2.02 / 10.0 | RL (mm) | 5 | shot mode | multishot
REC voxel MPS (mm) | 1.33 / 1.33 / 5.00 | Recon voxel size (mm) | 1.333333 | EPI factor | 5
Scan percentage (%) | 98.83721 | Fold-over suppression | no | epi direction | Y
Act. WFS (pix) / BW (Hz) | 2.182 / 199.1 | Slice oversampling | user defined | Echoes | 1
BW in EPI freq. dir. (Hz) | 1757.3 | oversample factor | 1.5 | partial echo | no
Min. WFS (pix) / Max. BW (Hz) | 2.181 / 199.2 | RF select. FOS | no | shifted echo | no
Min. TR/TE (ms) | 45 / 6.4 | Reconstruction matrix | 192 | TE | shortest
SAR / local torso | < 76 % | SENSE | yes | Flip angle (deg) | 6
Whole body / level | < 1.3 W/kg / normal | P reduction (FH) | 1.5 | TR | user defined
SED | < 0.1 kJ/kg | S reduction (RL) | 1 | (ms) | 48
B1+rms / Coil Power | 1.99 uT / 73 % | k-t BLAST | no | Halfscan | no
Max B1+rms | 1.99 uT | Overcontiguous slices | yes | Water-fat shift | maximum
PNS / level | 72 % / normal | Stacks | 1 | RF Shims | smart
dB/dt | 53.4 T/s | slices | 10 | Shim | IB-smart
Sound Pressure Level (dB) | 18.98915 | slice orientation | sagittal | mDIXON | no
max d¦B¦/dt (epi) [Tesla/second] | 53.4 (51.4) | fold-over direction | FH | Fat suppression | ProSet
MOTION |  | fat shift direction | F | pulse type | 1331
Cardiac synchronization | no | Multi-chunk | no | Water suppression | no
Heart rate > 250 bpm | no | PlanAlign | no | MTC | +qMT/CEST
Respiratory compensation | gate | REST slabs | 0 | +Duration | 20
Resp. cycle length (ms) | 2000 | Catheter tracking | no | +B1 Mode | arrayed (dyn. loop)
Gate delay (ms) | 350 | Interactive positioning | no | +Amps | 500, 500, 500, 500, 800, 800, 800, 800, 800, 850, 1E-10, 850, 850, 850, 850, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10, 1E-10
width (ms) | 1200 | External control | no | +B1 Units | flip angle (deg)
Navigator respiratory comp | no | Allow table movement | no | +Pulses/TR | 1
Flow compensation | yes | OFFC/ANG |  | +Offset Mode | user-defined
fMRI echo stabilisation | no | Stacks | 1 | +Offsets | 1000, 2000, 4000, 8000, 1000, 2000, 4000, 8000, 100000, 16000, 100000, 2000, 4000, 8000, 16000, 100000, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0, 0
NSA | 1 | Stack Offc. AP (P=+mm) | -25 | +Offset Units | Hz
DYN/ANG |  | RL (L=+mm) | -100 | +RF Shape | sg_100_100_0
Angio / Contrast enh. | no | FH (H=+mm) | 25 | Research prepulse | no
Quantitative flow | no | Ang. AP (deg) | 0 | Diffusion mode | no
Manual start | no | RL (deg) | 0 | Elastography mode | no
+Abuse dynamic loop | no | FH (deg) | 0 | Multi-transmit | yes
Dynamic study | individual | Free rotatable | no | Transmit channels | both
dyn scans | 9 |  |  | SAR mode | high
recon multiplier | 1 |  |  | B1 mode | default
dyn scan times | shortest |  |  | SAR Patient data | auto
fov time mode | default |  |  | PNS mode | high
dummy scans | 0 |  |  | Gradient mode | default
immediate subtraction | no |  |  | SofTone mode | no
fast next scan | no
synch. ext. device | no
MTC | yes
dyn stabilization | no
prospect. motion corr. | no
Keyhole | no
Arterial Spin labeling | no
POST/PROC
Preparation phases | +sameprep
Interactive F0 | no
SmartPlan survey | no
+Whole Body Fat-Water | no
B0 field map | no
B1 field map | no
MIP/MPR | no
SWIp | no
Images | M, no, no, no
Autoview image | M
Calculated images | no, no, no, no
Reference tissue | Grey matter
EPI 2D phase correction | no
Recon compression | No
Preset window contrast | soft
Reconstruction mode | real time
reuse memory | no
Save raw data | no
Hardcopy protocol | no
Image filter | system default
Geometry correction | default
Elliptical k-space shutter | default

## Variable Flip Angle
multi-flip_T1-map 01:06.9
INFO PAGE |  | GEOMETRY |  | CONTRAST
Total scan duration | 01:06.9 | Nucleus | H1 | Scan type | Imaging
Rel. SNR | 1 | Uniformity | Body-tuned | Scan mode | 3D
Act. TR/TE (ms) | 7.9 / 4.6 | FOV FH (mm) | 256 | technique | FFE
Dyn. scan time | 00:06.7 | AP (mm) | 256 | loop order | zy_order
Time to k0 | 00:03.3 | RL (mm) | 50 | Contrast enhancement | T1
ACQ matrix M x P | 192 x 192 | Voxel size FH (mm) | 1.333333 | Acquisition mode | cartesian
ACQ voxel MPS (mm) | 1.33 / 1.33 / 5.00 | AP (mm) | 1.333333 | Fast Imaging mode | none
REC voxel MPS (mm) | 1.33 / 1.33 / 5.00 | RL (mm) | 5 | 3D non-selective | no
Scan percentage (%) | 100 | Recon voxel size (mm) | 1.333333 | Echoes | 1
Act. WFS (pix) / BW (Hz) | 1.001 / 434.0 | Fold-over suppression | no | partial echo | yes
Min. WFS (pix) / Max. BW (Hz) | 0.639 / 680.2 | Slice oversampling | default | shifted echo | no
Min. TR/TE (ms) | 7.9 / 2.5 | RF select. FOS | no | TE | in-phase
SAR / local torso | < 38 % | Reconstruction matrix | 192 | (ms) | 4.603275
Whole body / level | < 0.7 W/kg / normal | SENSE | yes | Flip angle (deg) | 20
SED | 0.0 kJ/kg | P reduction (AP) | 1 | TR | shortest
B1+rms / Coil Power | 1.42 uT / 37 % | S reduction (RL) | 2 | Halfscan | no
Max B1+rms | 1.42 uT | k-t BLAST | no | Water-fat shift | user defined
PNS / level | 12 % / normal | Overcontiguous slices | no | (pixels) | 1
dB/dt | 6.4 T/s | Stacks | 1 | RF Shims | smart
Sound Pressure Level (dB) | 1.483045 | slices | 10 | Shim | IB-smart
max d¦B¦/dt (epi) [Tesla/second] | 6.4 (0.0) | slice orientation | sagittal | mDIXON | no
MOTION |  | fold-over direction | AP | Fat suppression | no
Cardiac synchronization | no | fat shift direction | F | Water suppression | no
Heart rate > 250 bpm | no | Multi-chunk | no | MTC | no
Respiratory compensation | no | PlanAlign | no | Research prepulse | no
Navigator respiratory comp | no | REST slabs | 0 | Diffusion mode | no
Flow compensation | no | Catheter tracking | no | Elastography mode | no
fMRI echo stabilisation | no | Interactive positioning | no | Multi-transmit | yes
NSA | 1 | External control | no | Transmit channels | both
DYN/ANG |  | Allow table movement | no | SAR mode | high
Angio / Contrast enh. | no | OFFC/ANG |  | B1 mode | default
Quantitative flow | no | Stacks | 1 | SAR Patient data | auto
Manual start | no | Stack Offc. AP (P=+mm) | -25 | PNS mode | low
+Abuse dynamic loop | flip angle diminish | RL (L=+mm) | -100 | Gradient mode | default
Dynamic study | individual | FH (H=+mm) | 25 | SofTone mode | yes
dyn scans | 10 | Ang. AP (deg) | 0
recon multiplier | 1 | RL (deg) | 0
dyn scan times | shortest | FH (deg) | 0
fov time mode | default | Free rotatable | no
dummy scans | 0
immediate subtraction | no
fast next scan | no
synch. ext. device | no
dyn stabilization | no
prospect. motion corr. | no
Keyhole | no
Arterial Spin labeling | no
POST/PROC
Preparation phases | +sameprep
Interactive F0 | no
SmartPlan survey | no
+Whole Body Fat-Water | no
B0 field map | no
B1 field map | no
MIP/MPR | no
SWIp | no
Images | M, no, no, no
Autoview image | M
Calculated images | no, no, no, no
Reference tissue | Breast
Recon compression | No
Preset window contrast | soft
Reconstruction mode | real time
reuse memory | no
Save raw data | no
Hardcopy protocol | no
Image filter | system default
Geometry correction | default
Elliptical k-space shutter | default

## DCE
dynamic_smartTX_6deg 06:07.8
INFO PAGE |  | GEOMETRY |  | CONTRAST
Total scan duration | 06:07.8 | Nucleus | H1 | Scan type | Imaging
Rel. SNR | 1 | Uniformity | Body-tuned | Scan mode | 3D
Act. TR/TE (ms) | 7.9 / 4.6 | FOV FH (mm) | 256 | technique | FFE
Dyn. scan time | 00:06.7 | AP (mm) | 256 | loop order | zy_order
Time to k0 | 00:03.3 | RL (mm) | 50 | Contrast enhancement | T1
ACQ matrix M x P | 192 x 192 | Voxel size FH (mm) | 1.333333 | Acquisition mode | cartesian
ACQ voxel MPS (mm) | 1.33 / 1.33 / 5.00 | AP (mm) | 1.333333 | Fast Imaging mode | none
REC voxel MPS (mm) | 1.33 / 1.33 / 5.00 | RL (mm) | 5 | 3D non-selective | no
Scan percentage (%) | 100 | Recon voxel size (mm) | 1.333333 | Echoes | 1
Act. WFS (pix) / BW (Hz) | 1.001 / 434.0 | Fold-over suppression | no | partial echo | yes
Min. WFS (pix) / Max. BW (Hz) | 0.639 / 680.2 | Slice oversampling | default | shifted echo | no
Min. TR/TE (ms) | 7.9 / 2.5 | RF select. FOS | no | TE | in-phase
SAR / local torso | < 38 % | Reconstruction matrix | 192 | (ms) | 4.603275
Whole body / level | < 0.7 W/kg / normal | SENSE | yes | Flip angle (deg) | 20
SED | < 0.2 kJ/kg | P reduction (AP) | 1 | TR | shortest
B1+rms / Coil Power | 1.42 uT / 37 % | S reduction (RL) | 2 | Halfscan | no
Max B1+rms | 1.42 uT | k-t BLAST | no | Water-fat shift | user defined
PNS / level | 12 % / normal | Overcontiguous slices | no | (pixels) | 1
dB/dt | 6.4 T/s | Stacks | 1 | RF Shims | smart
Sound Pressure Level (dB) | 1.483045 | slices | 10 | Shim | IB-smart
max d¦B¦/dt (epi) [Tesla/second] | 6.4 (0.0) | slice orientation | sagittal | mDIXON | no
MOTION |  | fold-over direction | AP | Fat suppression | no
Cardiac synchronization | no | fat shift direction | F | Water suppression | no
Heart rate > 250 bpm | no | Multi-chunk | no | MTC | no
Respiratory compensation | no | PlanAlign | no | Research prepulse | no
Navigator respiratory comp | no | REST slabs | 0 | Diffusion mode | no
Flow compensation | no | Catheter tracking | no | Elastography mode | no
fMRI echo stabilisation | no | Interactive positioning | no | Multi-transmit | yes
NSA | 1 | External control | no | Transmit channels | both
DYN/ANG |  | Allow table movement | no | SAR mode | high
Angio / Contrast enh. | no | OFFC/ANG |  | B1 mode | default
Quantitative flow | no | Stacks | 1 | SAR Patient data | auto
Manual start | yes | Stack Offc. AP (P=+mm) | -25 | PNS mode | low
+Abuse dynamic loop | no | RL (L=+mm) | -100 | Gradient mode | default
Dynamic study | individual | FH (H=+mm) | 25 | SofTone mode | yes
dyn scans | 55 | Ang. AP (deg) | 0
recon multiplier | 1 | RL (deg) | 0
dyn scan times | user defined | FH (deg) | 0
(mm:ss) | manual (00:00.0), shortest (00:06.7), shortest (00:13.4), manual (00:20.1), shortest (00:26.8), shortest (00:33.5), shortest (00:40.2), shortest (00:46.9), shortest (00:53.6), shortest (01:00.2), shortest (01:06.9), shortest (01:13.6), shortest (01:20.3), shortest (01:27.0), shortest (01:33.7), shortest (01:40.4), shortest (01:47.0), shortest (01:53.7), shortest (02:00.4), shortest (02:07.1), shortest (02:13.8), shortest (02:20.5), shortest (02:27.2), shortest (02:33.8), shortest (02:40.5), shortest (02:47.2), shortest (02:53.9), shortest (03:00.6), shortest (03:07.3), shortest (03:14.0), shortest (03:20.6), shortest (03:27.3), shortest (03:34.0), shortest (03:40.7), shortest (03:47.4), shortest (03:54.1), shortest (04:00.8), shortest (04:07.4), shortest (04:14.1), shortest (04:20.8), shortest (04:27.5), shortest (04:34.2), shortest (04:40.9), shortest (04:47.6), shortest (04:54.2), shortest (05:00.9), shortest (05:07.6), shortest (05:14.3), shortest (05:21.0), shortest (05:27.7), shortest (05:34.4), shortest (05:41.0), shortest (05:47.7), shortest (05:54.4), shortest (06:01.1), _, _, _, _, _, _, _, _, _ | Free rotatable | no
fov time mode | default
dummy scans | 0
immediate subtraction | no
fast next scan | no
synch. ext. device | no
dyn stabilization | no
prospect. motion corr. | no
Keyhole | no
Arterial Spin labeling | no
POST/PROC
Preparation phases | +sameprep
Interactive F0 | no
SmartPlan survey | no
+Whole Body Fat-Water | no
B0 field map | no
B1 field map | no
MIP/MPR | no
SWIp | no
Images | M, no, no, no
Autoview image | M
Calculated images | no, no, no, no
Reference tissue | Breast
Recon compression | No
Preset window contrast | soft
Reconstruction mode | real time
reuse memory | no
Save raw data | no
Hardcopy protocol | no
Image filter | system default
Geometry correction | default
Elliptical k-space shutter | default

## T1w
THRIVE SENSE 02:11.8
INFO PAGE |  | GEOMETRY |  | CONTRAST
Total scan duration | 02:11.8 | Nucleus | H1 | Scan type | Imaging
Rel. SNR | 1 | Uniformity | Body-tuned | Scan mode | 3D
Act. TR/TE (ms) | 6.2 / 3.1 | FOV FH (mm) | 256 | technique | FFE
ACQ matrix M x P | 512 x 512 | AP (mm) | 256 | Contrast enhancement | T1
ACQ voxel MPS (mm) | 0.50 / 0.50 / 2.00 | RL (mm) | 100 | Acquisition mode | cartesian
REC voxel MPS (mm) | 0.50 / 0.50 / 1.00 | Voxel size FH (mm) | 0.5 | Fast Imaging mode | TFE
Scan percentage (%) | 100 | AP (mm) | 0.5 | 3D non-selective | no
TFE shots | 205 | RL (mm) | 1 | shot mode | multishot
TFE dur. shot / acq (ms) | 636.8 / 495.4 | Recon voxel size (mm) | 0.5 | TFE factor | 80
TFE shot interval (ms) | 636.8121 | Fold-over suppression | rest | 3D free factor | no
Act. WFS (pix) / BW (Hz) | 1.119 / 388.4 | slab thickness | half FOV | startup echoes | user defined
Min. WFS (pix) / Max. BW (Hz) | 1.119 / 388.4 | Slice oversampling | default | (number) | 0
SPAIR offset act./default (Hz) | 220 [220] | RF select. FOS | no | +TFE follow-up echoes | 0
SAR / local torso | < 30 % | Reconstruction matrix | 512 | shot interval | shortest
Whole body / level | < 0.5 W/kg / normal | SENSE | yes | profile order | low_high
SED | < 0.1 kJ/kg | P reduction (AP) | 1 | turbo direction | radial
B1+rms / Coil Power | 1.25 uT / 28 % | S reduction (RL) | 2 | CENTRA (spiral) | no
Max B1+rms | 1.25 uT | k-t BLAST | no | Echoes | 1
PNS / level | 58 % / normal | Overcontiguous slices | yes | partial echo | no
dB/dt | 59.0 T/s | Stacks | 1 | shifted echo | no
Sound Pressure Level (dB) | 16.48793 | slices | 100 | TE | shortest
max d¦B¦/dt (epi) [Tesla/second] | 59.0 (0.0) | slice orientation | sagittal | Flip angle (deg) | 10
MOTION |  | fold-over direction | AP | TR | shortest
Cardiac synchronization | no | fat shift direction | F | Halfscan | no
Heart rate > 250 bpm | no | Multi-chunk | no | Water-fat shift | minimum
Respiratory compensation | no | PlanAlign | no | RF Shims | smart
Navigator respiratory comp | no | REST slabs | 0 | Shim | IB-smart
Flow compensation | no | Catheter tracking | no | mDIXON | no
fMRI echo stabilisation | no | Interactive positioning | no | Fat suppression | SPAIR
Motion smoothing | no | External control | no | power | 2
NSA | 1 | Allow table movement | no | inversion delay | auto
DYN/ANG |  | OFFC/ANG |  | suppr. level | strong
Angio / Contrast enh. | no | Stacks | 1 | frequency offset | user defined
Quantitative flow | no | Stack Offc. AP (P=+mm) | -25 | offset (Hz) | 220
CENTRA | no | RL (L=+mm) | -100 | Water suppression | no
Manual start | no | FH (H=+mm) | 25 | TFE prepulse | no
+Abuse dynamic loop | no | Ang. AP (deg) | 0 | MTC | no
Dynamic study | no | RL (deg) | 0 | T2prep | no
Arterial Spin labeling | no | FH (deg) | 0 | Research prepulse | no
POST/PROC |  | Free rotatable | no | Diffusion mode | no
Preparation phases | auto |  |  | Elastography mode | no
Interactive F0 | no |  |  | Multi-transmit | yes
SmartPlan survey | no |  |  | Transmit channels | both
+Whole Body Fat-Water | no |  |  | SAR mode | low
B0 field map | no |  |  | B1 mode | default
B1 field map | no |  |  | SAR Patient data | auto
MIP/MPR | no |  |  | PNS mode | low
SWIp | no |  |  | Gradient mode | default
Images | M, no, no, no |  |  | SofTone mode | no
Autoview image | no
Calculated images | no, no, no, no
Reference tissue | Breast
Recon compression | No
Preset window contrast | soft
Reconstruction mode | real time
Save raw data | no
Hardcopy protocol | no
Image filter | system default
Geometry correction | default
Elliptical k-space shutter | default